## Slide 1: Improving Electron Density Calibration Techniques for the DIII-D Thomson Scattering Diagnostic
Fabio Conti 1,*
Fenton Glass 1
Richard Carr 1
Adam G. McLean 2
1General Atomics
San Diego, Calif., USA
2Lawrence Livermore National Laboratory
Livermore, Calif., USA
*contif@fusion.gat.com

## Slide 2: Background & Motivation
Accurate electron density (ne) profiles are crucial:
  plasma control, performance evaluation, study confinement, transport, edge physics, & exhaust
  require absolute calibration of system response

## Slide 3: Background & Motivation
Accurate electron density (ne) profiles are crucial:
  plasma control, performance evaluation, study confinement, transport, edge physics, & exhaust
  require absolute calibration of system response
Rayleigh scattering-based calibration of DIII-D Thomson scattering[1] vulnerable to stray laser light contamination
Locations with high stray light (e.g., in the divertor) cannot be directly calibrated using Rayleigh scattering

## Slide 4: Background & Motivation
Accurate electron density (ne) profiles are crucial:
  plasma control, performance evaluation, study confinement, transport, edge physics, & exhaust
  require absolute calibration of system response
Rayleigh scattering-based calibration of DIII-D Thomson scattering[1] vulnerable to stray laser light contamination
Locations with high stray light (e.g., in the divertor) cannot be directly calibrated using Rayleigh scattering
Rotational Raman scattering from diatomic gas shifts scattered wavelengths away from laser wavelength[2]
[1] B. Bray et al., RSI 72, 1115 (2001) [2] J. Howard, J. Phys. D 12 (1979)

## Slide 5: Rayleigh scattering
System response measured at various gas fill pressures
Signal normalized to laser energy
Slope of linear fit provides calibration factor
Single spectral channel only (1064nm)
Well-known cross section

## Slide 6: Raman scattering
Sum of Raman lines weighted by filter transmission

## Slide 7: Raman scattering
Sum of Raman lines weighted by filter transmission
Usability of spectral channels:
  1047nm, 1058nm reliably good signals
  1062nm, 1067nm possible Rayleigh bleed-through
  1022nm low Raman signal

## Slide 8: Raman scattering
Sum of Raman lines weighted by filter transmission
Usability of spectral channels:
  1047nm, 1058nm reliably good signals
  1062nm, 1067nm possible Rayleigh bleed-through
  1022nm low Raman signal
100-1000x lower signal (per filter) vs Rayleigh

## Slide 9: Rayleigh vs Raman
56 scattering volumes (42 vert. + 14 hor.) over (6 vert. + 1 hor.) lasers → 266 coefficients have both Rayleigh and Raman values

## Slide 10: Rayleigh vs Raman
56 scattering volumes (42 vert. + 14 hor.) over (6 vert. + 1 hor.) lasers → 266 coefficients have both Rayleigh and Raman values
Geometry factor Φ should be the same for Rayleigh / Raman (& Thomson)
For each spectral channel:

## Slide 11: Rayleigh vs Raman
56 scattering volumes (42 vert. + 14 hor.) over (6 vert. + 1 hor.) lasers → 266 coefficients have both Rayleigh and Raman values
Geometry factor Φ should be the same for Rayleigh / Raman (& Thomson)
For each spectral channel:
ratio of detector gains (DC spectral calibrations)
relative transmission
of collection optics + filter
ratio of detector
signal to laser
energy
cross-section

## Slide 12: Rayleigh vs Raman
Geometry factor Φ should be the same for Rayleigh / Raman (& Thomson)
DIII-D Thomson system produces lower Raman sensitivity than Rayleigh: → gives ~25% higher ne for same Thomson signal
Notes: Mention bulk comparison with CO2?

## Slide 13: Rayleigh vs Raman
Raman σ ∝ γ2 where γ = molecular polarizability anisotropy[3]
Manually adjust value to improve agreement γ2 = 0.51 → 0.40 x10-48 cm6
Still use Rayleigh where possible until discrepancy is resolved
Geometry factor Φ should be the same for Rayleigh / Raman (& Thomson)
DIII-D Thomson system produces lower Raman sensitivity than Rayleigh: → gives ~25% higher ne for same Thomson signal
[3] B. P. LeBlanc, RSI 79, 10E737 (2008)
Notes: Mention bulk comparison with CO2?

## Slide 14
Ratio ΓRaman / ΓRayleigh has ~20% point-to-point variability
Rayleigh vs Raman
[Raman value is mean of 1058 & 1047 nm channels]

## Slide 15
Ratio ΓRaman / ΓRayleigh has ~20% point-to-point variability
Possible causes:
  measurement accuracy of detector Pulse / DC gain → whole system pulsed calibration
  accuracy of optics spectral calibrations: angle-of-incidence, coating, aging, irradiation → in situ (re)calibrations
  Residual Rayleigh cross-filter bleed-through during high-pressure Raman? → measure
Rayleigh vs Raman
[Raman value is mean of 1058 & 1047 nm channels]

## Slide 16: Rayleigh vs Raman
Neither 1047nm nor 1058nm factor alone is better matched to Rayleigh
If ‘good’ is within 10% of Rayleigh:
Good channel(s) | Frequency
1047nm only | 28%
1058nm only | 20%
Both | 28%
Neither | 23%
Mean of channels | 58%

## Slide 17: Rayleigh vs Raman
Neither 1047nm nor 1058nm factor alone is better matched to Rayleigh
If ‘good’ is within 10% of Rayleigh:
Good channel(s) | Frequency
1047nm only | 28%
1058nm only | 20%
Both | 28%
Neither | 23%
Mean of channels | 58%
Where both are available, Raman is not clearly better than Rayleigh at DIII-D

## Slide 18: Cross-calibration with ECE improves ne calibration and profiles
Dedicated shots with gradual ramp-up of ne and Ip to send ECE channels into cutoff
[4] K. Zhurovich et al., RSI 76, 053506 (2005) [5] M. Brookman et al., RSI 87, 11E517 (2016)
Fit signal extinction: known ECE fn and B(R) gives ne(Rn, tn) [4, 5]
Mapping + interpolation in ψ space to get correction factors for Rayleigh/Raman
Can introduce errors ∝ 𝛻ψ

## Slide 19: Cross-calibration with ECE improves ne calibration and profiles
DIII-D shot 205686, t = 2750 ms
Use ECE correction to Rayleigh/Raman in plasma core & edge (53/56 volumes)

## Slide 20: Density calibration is more challenging in divertor
Ultra-narrowband filters reduce cross-filter stray light contamination
  Rayleigh calibration not possible
Raman calibration used with adjusted σ
Possible cross-checks / corrections being evaluated
  ECE-based
  Pe based

## Slide 21: Calibration of Divertor Density – ECE and Pe checks
ECE: expanded plasma shape, divertor TS locations inside separatrix
  ECE mapped from mid-plane (steep 𝛻Ψ) to expanded flux in divertor
Divertor scattering volumes
ECE mapped from midplane

## Slide 22: Calibration of Divertor Density – ECE and Pe checks
ECE: expanded plasma shape, divertor TS locations inside separatrix
  ECE mapped from mid-plane (steep 𝛻Ψ) to expanded flux in divertor.
Assuming Pe balance along field lines for independent check
  Radial sweep of strike point, L-mode steady state. Each volume sees varying Ψ(t).
  Corrections relative to a reference volume
Divertor scattering volumes
ECE mapped from midplane

## Slide 23: Calibration of Divertor Density – ECE and Pe checks
Independent ECE and Pe methods complement each other
  agreement within ~10% in most locations
  Can use Pe near wall and ECE in top-most volumes
Notes: Due to sweep, getting overlap between adjacent points through nearby psi.
More complicated physics for dl00 and dl01 near the tile.
CHECK ECE VS PE values (should match at -106 cm)
ADD ERROR BARS
Show DTS profiles?

## Slide 24: Calibration of Divertor Density – ECE and Pe checks
Independent ECE and Pe methods complement each other
  agreement within ~10% in most locations
  Can use Pe near wall and ECE in top-most volumes
  To be implemented in new data for smoother profiles
Notes: Due to sweep, getting overlap between adjacent points through nearby psi.
More complicated physics for dl00 and dl01 near the tile.
CHECK ECE VS PE values (should match at -106 cm)
ADD ERROR BARS
Show DTS profiles?

## Slide 25: Conclusion & Further Development
Raman and Rayleigh scattering used to calibrate DIII-D TS for ne
Raman underestimates sensitivity: systematic error or adjust γ Where both are available, Raman is not clearly better than Rayleigh
ECE cutoff used to refine Rayleigh calibrations / ne profiles in plasma core & edge beam paths
Divertor regions use adjusted-γ Raman; ECE and P-balance used to cross-check results, combining the two seems to work well
Calibrating diagnostics in multiple ways – understand discrepancies
  point-to-point and spectral channel-to-channel variation
  consistency between techniques
  improve ECE and P-balance correction methods
Calibration process improvements: whole system pulsed calibration, measure Rayleigh cross-filter bleed-through
Notes: Summarize what factors are used where. Circle back to motivation of precise profiles
Do we need Rayleigh/Raman as baseline? Yes, since cross-calibrations need critical analysis and are not applied everywhere.
Measuring the same thing multiple ways. Understanding why they don’t agree (“discrepancies”)

## Slide 26
THANK YOU

## Slide 28: Comparison of Raman spectral channels

## Slide 29: ECE cutoff
Ip
ne
Te

## Slide 30: Divertor strike point sweep